# Metering Installation Error Report

# Accuracy

The accuracy of the metering installation (combined meter and transformers) has been calculated for 3 annual load profiles.

Load Profile		Error		Uncertainty	Error-Uncert	Error+Uncert

1			2.76 %		0.74 %		2.01 %		3.50 %

2			1.41 %		0.40 %		1.00 %		1.81 %

3			0.58 %		0.21 %		0.36 %		0.79 %

The uncertainty is the expanded uncertainty calculated at a 95 % level of confidence. For each load profile the same variation in temperature and network conditions was assumed. Sensitivity coefficients published for an IEC class 0.2 meter were incorporated in the uncertainty calculation.

The uncertainty includes a contribution due to the temperature of the metering installation differing from the temperature at which the components were calibrated. An average value of 15.0 °C with an expanded uncertainty of 1.0 °C is assumed for a twelve-month period.

# Annual Load Profiles

Annual load profiles used for this error calculation. The graph is normalised to give the relative amount of energy at each combination of current and phase angle.

# Error Plots

Installation error over the phase and current values for the load profiles

# Uncertainty by Component

Uncertainty contribution given by component as a percentage of total variance. This is calculated for each load profile. This may help with any decisions about which components might be considered for an upgrade

Load Profile		Meter		CT	VT

1			54 %		15 %	31 %

2			50 %		10 %	40 %

3			60 %		15 %	25 %

# Component Calibration

For each component the calibration points are plotted on a graph with uncertainty bars. Fitted curves are shown with an upper and lower curve indicating the uncertainty  of the fit. The coefficients of these curves are used to calculate the metering installation error.

# Current Transformer

# Voltage Transformer

# Meter

# Installation Maximum Error Contour

# Additional Information

## Text Output From the Calculation Process

Meter data

Data scatter greater than expected, a poor fit.

CT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT phase data

Good fit.

VT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

VT phase data

Fit has lower uncertainty than expected from calibration data, so type B component added.

Load phase limited to 89.9 degrees maximum.

For profile in  C:\Users\k.jones\OneDrive - Callaghan Innovation\KJ\PycharmProjects\MIEcalc3\Projects

\mie_temp\load.txt

Max current = 86.65 %, Min current = 20.35 %, Max phase = 89.90 degrees, Min phase = 61.93 degrees

Load phase limited to 89.9 degrees maximum.

For profile in  C:\Users\k.jones\OneDrive - Callaghan Innovation\KJ\PycharmProjects\MIEcalc3\Projects

\mie_temp\load1.txt

Max current = 16.78 %, Min current = 2.41 %, Max phase = 89.90 degrees, Min phase = -11.13 degrees

Load error, low or negative current.

Load phase limited to 89.9 degrees maximum.

For profile in  C:\Users\k.jones\OneDrive - Callaghan Innovation\KJ\PycharmProjects\MIEcalc3\Projects

\mie_temp\load2.txt

Max current = 193.16 %, Min current = 0.10 %, Max phase = 89.90 degrees, Min phase = -19.92 degrees

start site error

Load Profile # 1

Budget for overall error by coefficients

error = 2.757, standard uncertainty = 0.361, degrees of freedom = 25.8

a2 meter: 0.302478

VT calibration burden: 0.29453

a3 meter: 0.293324

a2 CT_p: 0.181719

b0 VT_p: 0.132538

a0 CT_p: 0.114243

a1 CT_p: 0.0513759

Meter temperature coefficient: 0.040078

Site VT burden: 0.029453

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

Meter field dependence: 0.0176777

Site EM field: 0.0176777

CT calibration burden: 0.0174925

b0 VT_e: 0.0163636

a0 meter: 0.015926

b0 CT_e: 0.0140929

a0 VT_p: 0.0126774

a1 VT_p: 0.0125524

Meter voltage dependence: 0.0106066

Meter annual drift: 0.01

Site voltage: 0.00860223

a2 CT_e: 0.00756254

a0 VT_e: 0.00664262

a1 VT_e: 0.00657718

Site frequency: 0.00647566

Meter frequency dependence: 0.0053033

Site temperature: 0.00510347

VT temperature coefficient of phase: 0.00471247

a1 meter: 0.00455804

a0 CT_e: 0.00405622

VT calibration temperature: 0.0029453

a1 CT_e: 0.00247547

Meter calibration temperature: 0.00176777

Site CT burden: 0.00174925

CT temperature coefficient of phase: 0.00071668

zero burden factor: 0.000394415

CT temperature coefficient of error: 9.20302E-05

CT calibration temperature: 7.80812E-05

start site error

Load Profile # 2

Budget for overall error by coefficients

error = 1.405, standard uncertainty = 0.199, degrees of freedom = 34.8

a2 meter: 0.152815

VT calibration burden: 0.1488

a3 meter: 0.0871667

b0 VT_p: 0.0669601

a0 CT_p: 0.0577172

a2 CT_p: 0.0540012

Meter temperature coefficient: 0.040078

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

Meter field dependence: 0.0176777

Site EM field: 0.0176777

CT calibration burden: 0.0175606

b0 VT_e: 0.0163636

a0 meter: 0.015926

Site VT burden: 0.01488

b0 CT_e: 0.0140929

Meter voltage dependence: 0.0106066

Site voltage: 0.0103856

Meter annual drift: 0.01

a0 VT_e: 0.00664262

a1 VT_e: 0.00657718

a0 VT_p: 0.00640476

a1 VT_p: 0.00634166

Site frequency: 0.0059272

Meter frequency dependence: 0.0053033

a1 CT_p: 0.00524587

a2 CT_e: 0.00442777

a0 CT_e: 0.00405622

Site temperature: 0.00366844

VT temperature coefficient of phase: 0.0023808

Meter calibration temperature: 0.00176777

Site CT burden: 0.00175606

VT calibration temperature: 0.001488

a1 meter: 0.000910132

CT temperature coefficient of phase: 0.000791389

zero burden factor: 0.00056192

a1 CT_e: 0.000494293

CT temperature coefficient of error: 0.000131115

CT calibration temperature: 8.25343E-05

start site error

Load Profile # 3

Budget for overall error by coefficients

error = 0.575, standard uncertainty = 0.106, degrees of freedom = 62.7

a3 meter: 0.072325

a2 meter: 0.0673792

VT calibration burden: 0.0656087

a2 CT_p: 0.0448066

Meter temperature coefficient: 0.040078

b0 VT_p: 0.029524

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

a0 CT_p: 0.0254486

a1 CT_p: 0.0184513

Meter field dependence: 0.0176777

Site EM field: 0.0176777

b0 VT_e: 0.0163636

a0 meter: 0.015926

b0 CT_e: 0.0140929

Site voltage: 0.0114037

Meter voltage dependence: 0.0106066

Meter annual drift: 0.01

a2 CT_e: 0.00836194

a1 meter: 0.00734566

a0 VT_e: 0.00664262

a1 VT_e: 0.00657718

Site VT burden: 0.00656087

Meter frequency dependence: 0.0053033

Site frequency: 0.0050725

a0 CT_e: 0.00405622

a1 CT_e: 0.00398943

a0 VT_p: 0.00282398

a1 VT_p: 0.00279616

Site temperature: 0.00246011

Meter calibration temperature: 0.00176777

VT temperature coefficient of phase: 0.00104974

VT calibration temperature: 0.000656087

CT calibration burden: 0.000502787

zero burden factor: 0.000368504

CT temperature coefficient of phase: 0.000143998

CT temperature coefficient of error: 8.59842E-05

Site CT burden: 5.02787E-05

CT calibration temperature: 7.25173E-06

start error by point