# Metering Installation Error Report

# (Data source: explore_1.xlsx. Error calculated: Fri Mar 13 14:06:11 2020.)

The error calculated for the given load profile is +0.14 %.

The expanded uncertainty at an estimated 95% level of confidence is 0.26 %.

The uncertainty interval extends from -0.12 % to +0.39 %.

## Installation error over the phase and current ranges that match the load profile.

## Annual load profile used for this error calculation.

## Component calibration data with calculated fits over the calibration ranges.

### Current Transformer.

### Voltage Transformer.

### Meter.

# Additional Information

## Uncertainty Contribution Summary

Uncertainty contribution by component as a percentage of total variance, Meter:  34  %, CT:  43  %, VT:  23  %.

## Text Output From Calculation Process

Meter data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT phase data

Good fit.

VT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

VT phase data

Fit has lower uncertainty than expected from calibration data, so type B component added.

Max current = 97.28 %, Min current = 39.88 %, Max phase = 10.61  degrees, Min phase = -1.66  degrees

Budget for overall error by coefficients

0.14, u=0.13, df=43.4

a2 CT_e: 0.0821305

Meter temperature coefficient: 0.0788987

Site temperature: 0.0751666

a0 VT_e: 0.0515479

a1 VT_e: 0.05104

a0 CT_e: 0.0418284

a1 CT_e: 0.0362267

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

Site voltage: 0.0227537

Meter voltage dependence: 0.0212132

b0 meter: 0.0202546

b0 CT_e: 0.0183737

b0 VT_e: 0.0163636

Meter frequency dependence: 0.0132583

Site frequency: 0.0124963

CT calibration burden: 0.0101189

Site CT burden: 0.0101189

a0 meter: 0.00824024

a2 CT_p: 0.00723731

a1 meter: 0.0059021

a0 VT_p: 0.0045432

a1 VT_p: 0.00449843

a0 CT_p: 0.00368657

a1 CT_p: 0.00319006

a3 meter: 0.00162068

b0 VT_p: 0.0013822

a2 meter: 0.00133656

VT calibration burden: 0.00127822

Site VT burden: 0.00127822

CT temperature coefficient of error: 0.000181742

CT temperature coefficient of phase: 0.000129633

VT temperature coefficient of phase: 9.45169E-05

VT temperature coefficient of error: 5.51839E-05

CT calibration temperature: 3.95621E-05

VT calibration temperature: 4.99749E-06