# Metering Installation Error Report

# (Data source: explore_1.xlsx. Error calculated: Fri Mar 13 14:01:00 2020.)

The error calculated for the given load profile is +0.76 %.

The expanded uncertainty at an estimated 95% level of confidence is 0.21 %.

The uncertainty interval extends from +0.54 % to +0.97 %.

## Installation error over the phase and current ranges that match the load profile.

## Annual load profile used for this error calculation.

## Component calibration data with calculated fits over the calibration ranges.

### Current Transformer.

### Voltage Transformer.

### Meter.

# Additional Information

## Uncertainty Contribution Summary

Uncertainty contribution by component as a percentage of total variance, Meter:  15  %, CT:  45  %, VT:  40  %.

## Text Output From Calculation Process

Meter data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT phase data

Good fit.

VT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

VT phase data

Fit has lower uncertainty than expected from calibration data, so type B component added.

Load error, low or negative current

Load phase limited to 89.9 degrees maximum

Max current = 41.02 %, Min current = 0.10 %, Max phase = 89.90  degrees, Min phase = -26.95  degrees

Budget for overall error by coefficients

0.76, u=0.11, df=100.7

a2 CT_e: 0.0560681

a0 VT_e: 0.0515479

a1 VT_e: 0.05104

CT calibration burden: 0.0480323

Site CT burden: 0.0480323

a2 CT_p: 0.0425322

a0 CT_e: 0.0418284

a0 VT_p: 0.0395204

a1 VT_p: 0.0391311

Meter temperature coefficient: 0.0357071

a0 CT_p: 0.0320688

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

VT calibration burden: 0.0249255

Site VT burden: 0.0249255

Site voltage: 0.0224636

Meter voltage dependence: 0.0212132

b0 meter: 0.0202546

b0 CT_e: 0.0183737

b0 VT_e: 0.0163636

Site frequency: 0.0132676

Meter frequency dependence: 0.0132583

b0 VT_p: 0.0120235

a2 meter: 0.0116265

a1 CT_e: 0.0101311

Site temperature: 0.00991642

a3 meter: 0.00952438

a0 meter: 0.00824024

a1 CT_p: 0.00754207

a1 meter: 0.00165056

CT temperature coefficient of phase: 0.000767252

VT temperature coefficient of phase: 0.000364712

CT calibration temperature: 0.000187326

CT temperature coefficient of error: 0.000111611

VT calibration temperature: 9.72096E-05

VT temperature coefficient of error: 2.4479E-05