# Metering Installation Error Report

# (Data source: explore_1.xlsx. Error calculated: Fri Mar 13 13:58:08 2020.)

The error calculated for the given load profile is +0.81 %.

The expanded uncertainty at an estimated 95% level of confidence is 0.22 %.

The uncertainty interval extends from +0.59 % to +1.03 %.

## Installation error over the phase and current ranges that match the load profile.

## Annual load profile used for this error calculation.

## Component calibration data with calculated fits over the calibration ranges.

### Current Transformer.

### Voltage Transformer.

### Meter.

# Additional Information

## Uncertainty Contribution Summary

Uncertainty contribution by component as a percentage of total variance, Meter:  13  %, CT:  50  %, VT:  37  %.

## Text Output From Calculation Process

Meter data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT phase data

Good fit.

VT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

VT phase data

Fit has lower uncertainty than expected from calibration data, so type B component added.

Load phase limited to 89.9 degrees maximum

Max current = 54.83 %, Min current = 5.53 %, Max phase = 89.90  degrees, Min phase = -20.56  degrees

Budget for overall error by coefficients

0.81, u=0.11, df=98.4

a2 CT_e: 0.0666241

a2 CT_p: 0.0568575

a0 VT_e: 0.0515479

a1 VT_e: 0.05104

CT calibration burden: 0.0506465

Site CT burden: 0.0506465

a0 VT_p: 0.0440408

a1 VT_p: 0.0436068

a0 CT_e: 0.0418284

a0 CT_p: 0.0357368

Meter temperature coefficient: 0.0357071

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

VT calibration burden: 0.0279816

Site VT burden: 0.0279816

Site voltage: 0.0224261

Meter voltage dependence: 0.0212132

b0 meter: 0.0202546

b0 CT_e: 0.0183737

a1 CT_e: 0.0166465

b0 VT_e: 0.0163636

a1 CT_p: 0.0141734

b0 VT_p: 0.0133987

Site frequency: 0.0133681

Meter frequency dependence: 0.0132583

a2 meter: 0.0129563

a3 meter: 0.0127323

Site temperature: 0.0101376

a0 meter: 0.00824024

a1 meter: 0.00271207

CT temperature coefficient of phase: 0.000710804

VT temperature coefficient of phase: 0.000406428

CT calibration temperature: 0.000197522

VT calibration temperature: 0.000109128

VT temperature coefficient of error: 2.4479E-05

CT temperature coefficient of error: 1.94792E-05