# Metering Installation Error Report

# (Data source: explore_1.xlsx. Error calculated: Fri Mar 13 14:03:15 2020.)

The error calculated for the given load profile is +0.14 %.

The expanded uncertainty at an estimated 95% level of confidence is 0.15 %.

The uncertainty interval extends from -0.01 % to +0.29 %.

## Installation error over the phase and current ranges that match the load profile.

## Annual load profile used for this error calculation.

## Component calibration data with calculated fits over the calibration ranges.

### Current Transformer.

### Voltage Transformer.

### Meter.

# Additional Information

## Uncertainty Contribution Summary

Uncertainty contribution by component as a percentage of total variance, Meter:  17  %, CT:  54  %, VT:  29  %.

## Text Output From Calculation Process

Meter data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

CT phase data

Good fit.

VT ratio data

Fit has lower uncertainty than expected from calibration data, so type B component added.

VT phase data

Fit has lower uncertainty than expected from calibration data, so type B component added.

Max current = 100.79 %, Min current = 38.36 %, Max phase = 11.30  degrees, Min phase = -2.05  degrees

Budget for overall error by coefficients

0.136, u=0.075, df=106.5

a2 CT_e: 0.0821861

a0 VT_e: 0.0515479

a1 VT_e: 0.05104

a0 CT_e: 0.0418284

a1 CT_e: 0.0363591

Meter temperature coefficient: 0.0357071

Meter harmonic dependence: 0.0282843

Site harmonics: 0.0282843

Site voltage: 0.0227533

Meter voltage dependence: 0.0212132

b0 meter: 0.0202546

b0 CT_e: 0.0183737

b0 VT_e: 0.0163636

Meter frequency dependence: 0.0132583

Site frequency: 0.012499

CT calibration burden: 0.01017

Site CT burden: 0.01017

a0 meter: 0.00824024

Site temperature: 0.00751887

a2 CT_p: 0.00732883

a1 meter: 0.00592367

a0 VT_p: 0.00459336

a1 VT_p: 0.0045481

a0 CT_p: 0.00372727

a1 CT_p: 0.00325002

a3 meter: 0.00164117

b0 VT_p: 0.00139746

a2 meter: 0.00135132

VT calibration burden: 0.00131213

Site VT burden: 0.00131213

CT temperature coefficient of error: 8.07585E-05

CT temperature coefficient of phase: 5.80625E-05

VT temperature coefficient of phase: 4.23896E-05

CT calibration temperature: 3.96631E-05

VT temperature coefficient of error: 2.4479E-05

VT calibration temperature: 5.11731E-06